| Customer: | Ringline | Tested By: | D. Durbin |
| --- | --- | --- | --- |
| Job No: | SO10-198 | Temperature: | +25°C |
| Model No: | AUM-50M100M-DTOK-3R5MBW | Date: | 2/10/12 |
| Serial No: | PL10161 |  |  |

| Test Item No. | PARAMETERS | SPECIFIED VALUE | Test Results | qa/qc |
| --- | --- | --- | --- | --- |
| 1 | Frequency Range: | 50 – 100 MHz | 50 – 100 MHz |  |
| 2 | Frequency Accuracy: | ± 2.0 MHz Max @ 25°C | ± 0.01 MHz See Plot |  |
| 3 | Modulation Bandwidth (3dB): | 3.5 MHz Min | 3.5 MHz |  |
| 4 | Modulation Frequency Deviation BW: | 8 to 20 MHz Typ @ 2V P-P | > 8 MHz @ 2V P-P |  |
| 5 | RF Output Power: | +10 dBm Min | +11.74 dBm |  |
| 6 | Phase Noise: | -60 dBc / Hz Typ @ 100 kHz Offset | < -60 dBc / Hz @ 100 kHz Offset |  |
| 7 | Harmonics: | -20 dBc Typ -15 dBc Max | < -20 dBc |  |

| Test Item No. | PARAMETERS | SPECIFIED VALUE | Test Results | qa/qc |
| --- | --- | --- | --- | --- |
| 8 | Spurious: | -55 dBc Max | < -55 dBc |  |
| 9 | Linearized Frequency Tuning Step Size: | LSB 0.1 MHz Nominal for Full Band | 0.1 MHz LSB |  |
| 10 | DC Power: | +15V (±0.5V) -15 V (±0.5V) + 5V (±0.5V) +28V (±2.0V) | +15V (±0.5V) -15 V (±0.5V) + 5V (±0.5V) +28V (±2.0V) |  |

QA/QC Approval: ________________________________________________Date:_________